## P1 O2 General
| OXYGEN Hospital Survey- questions for Manager |  |  |  |  | Page 1
1 | Surveyor first, last name
| Hospital name | District |  | Region |  |  | Country | GPS
| Hospital Manager- first, last name |  |  |  | Phone:
| Hospital Technician first, last name |  |  |  | Phone:
2 | How many beds are available |  |  |  |  | Current | Planned
| a | Non-ICU beds
| b | Non-ICU beds designated for COVID-19
| c | ICU - adult
| d | ICU - pediatric
| e | ICU - neonatal
| f | Ancillary - added inside or proximal
3 | Autoclave and sterilizer equipment |  |  |  |  | Functional | Not Functional
| a | 40-60 litres
| b | >90 litres
| c | Describe other sterilizing capacities:
4 | How many days usually to have O2 cylinder refilled? | Enter X in one box: | <1 day | 1-2 days | 3-4 days | 5-7 days | >7 days
5 | Electrical power Average number of hours per day and average number of days per week the hospital has electrical power from grid. |  |  |  | Enter number 0 if none | Hours /day | Days /week | 5c Back-up sufficient for all concentrators?
|  | Back-up power: |  |  | Enter X in one box: | Yes | No | Yes | No*
| a | Are there solar panels + batteries off grid?
| b | Is there a functioning generator?
*If answer is no, go to Technician page 2- UPS, generators
6 | How many emergency transport vehicles/ambulances are there:
| a | With oxygen
| b | Without oxygen
7 | 7a If central piped oxygen, how many liters/day are available? * if central piped oxygen, fill out Technician p3 central oxygen survey | Enter liters/day, if piped O2, 0 if none | Liters/Day |  | 7b How is central piped oxygen produced for this facility? Enter X for all that apply*
[5 empty] | oxygen plant |  | liquid oxygen |  | cylinder manifold

## P2 O2 Equipment
O2 Survey P2: O2 Equipment
1. Oxygen-giving equipment/item, how many available on each ward? | Enter number, 0 if none If counting boxes or other units of items, enter numbers after multipling by supply per box. : | Emergency room | Pediatric ward | Adult ward | Operating room | Maternity ward | Intensive Care Unit | Neonatal Unit | IN STORAGE | TOTALS | How many are NOT working | Total working
1 | Nasal cannula | [8 empty] | 0 |  | 0
2 | Neonate/ Infant nasopharyngeal catheters | [8 empty] | 0 |  | 0
3 | Face masks - single use | [8 empty] | 0 |  | 0
4 | Face masks with reservoir bag | [8 empty] | 0 |  | 0
5 | Bag Valve mask-adult | [8 empty] | 0 |  | 0
6 | Bag Valve mask -child | [8 empty] | 0 |  | 0
7 | Intubation sets | [8 empty] | 0 |  | 0
8 | Oral airway | [8 empty] | 0 |  | 0
9 | Nasopharyngeal airway | [8 empty] | 0 |  | 0
10 | Crico-thyroidotomy tray | [8 empty] | 0 |  | 0
11 | Regulator/gauge- bullnose | [8 empty] | 0 |  | 0
12 | Regulator/gauge- pin index | [8 empty] | 0 |  | 0
13 | Flow splitter- # outlets | [8 empty] | 0 |  | 0
14 | Mechanical ventilator - adult | [8 empty] | 0 |  | 0
Also fill sheet 8 15 | Mechanical ventilator -child | [8 empty] | 0 |  | 0
Also fill sheet 8 16 | Mechanical ventilator - portable | [8 empty] | 0 |  | 0
17 | Bubble CPAP | [8 empty] | 0 |  | 0
Also fill sheet 8 18 | CPAP | [8 empty] | 0 |  | 0
Also fill sheet 8 19 | BiPAP | [8 empty] | 0 |  | 0
21 | O2 terminal wall units - with valve, pressure & flow regulator | [8 empty] | 0 |  | 0
22 | O2 terminal wall units without these valves and regulators | [8 empty] | 0 |  | 0
23 | Oxygen tubing- sets | [8 empty] | 0
24 | Patient monitor- integrated ECG | [8 empty] | 0 |  | 0
25 | Patient monitor - no ECG | [8 empty] | 0 |  | 0
26 | High flow nasal oxygen- child | [8 empty] | 0 |  | 0
27 | High flow nasal oxygen- adult | [8 empty] | 0 |  | 0
28 | Suction - manual | [8 empty] | 0 |  | 0
29 | Suction - electric | [8 empty] | 0 |  | 0
30 | Suction - central vacuum | [8 empty] | 0 |  | 0
31 | Humidifier | [8 empty] | 0 |  | 0
32 | Nebulizers | [8 empty] | 0 |  | 0
33 | Oximeter - finger tip | [8 empty] | 0 |  | 0
34 | Oximeter - hand held | [8 empty] | 0 |  | 0
35 | Oximeter - table top | [8 empty] | 0 |  | 0
36 | Probes for handheld, tabletop | [8 empty] | 0 |  | 0
37 | Blood gas analyzer | [8 empty] | 0 |  | 0
2. How many oxygen cylinders, of each size, in each ward? | For each row, enter # cylinders per ward, 0 if none: | Emergency room | Pediatric ward | Adult ward | Operating room | Maternity ward | Intensive Care Unit | Neonatal Unit | IN STORAGE | TOTAL cylinders | cost to refill per cylinder- USD $
a | Size C, 170 liters - pin index | [8 empty] | 0
b | Size D, 340 liters - pin index | [8 empty] | 0
c | Size E, 680 liters - pin index | [8 empty] | 0
d | Size F, 1360 liters - bullnose | [8 empty] | 0
e | Size G, 3400 liters - bullnose | [8 empty] | 0
f | Size H, 4100 liters - bullnose | [8 empty] | 0
g | Size J, 6800 liters - bullnose | [8 empty] | 0
| Total liters/ward | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | TOTAL cylinder liters
3. How many concentrators, by flow, in each ward? | For each row, enter # of concentrators for each ward, 0 if none: | Emergency room | Pediatric ward | Adult ward | Operating room | Maternity ward | Intensive Care Unit | Neonatal Unit | IN STORAGE | TOTALS | How many are NOT working | Total working
a | 3 liter flow | [8 empty] | 0 |  | 0
b | 5 liter flow | [8 empty] | 0 |  | 0
c | 8 liter flow | [8 empty] | 0 |  | 0
c | 10 liter flow | [8 empty] | 0 |  | 0
e | > 10 liter flow | [8 empty] | 0 |  | 0
| Total Liters/min/ward | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | TOTAL concentrator liters
| liters/day/ward | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | Emergency room | Pediatric ward | Adult ward | Operating room | Maternity ward | Intensive Care Unit | Neonatal Unit | IN STORAGE | TOTAL
TOTAL liters/ 24 hours If all concentrators functioning and full cylinders |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
TOTAL HOSPITAL SUPPLY IN LITERS PER 24H (CYLINDERS + CONCENTRATORS + CENTRAL) | [9 empty] | 0

## P3 Staff-Maintenance
OXYGEN HOSPITAL SURVEY Oxygen maintanence and training Page 3
Maintenance
1 | Oxygen Concentrators - Date last maintenance, schedule of mantenance
| 1.1 Maintenance by Hospital Staff - See cleaning notes in survey instructions |  |  | Enter date, X if weekly | Date of last | Weekly or before re-use
| a | Filter removed and cleaned
| b | Exterior cleaned
| 1.2 Maintenance by Technician |  |  | Enter X in one box: | Yes | No | Unk
| a | Are there trained technicians for maintenance?
| b | Is there inventory of spare parts for each type concentrator?
|  |  |  | Date of last | < 6 months | 6-12 months | >12 months | Only for Repairs
| c | When did Technician check compressor for flow and troubleshooting?
2 | Oxygen Concentrators - have clearance, no sunlight, and >= 1.5 m from heat sources? |  |  | Enter X in one box: | Yes | No | If no, describe:
3 | Humidifiers - Enter date of last maintenance and schedule. See cleaning notes in survey Instructions |  | Enter date, X if weekly or daily | Date of last | Daily | Weekly or before re-use
| a | Water changed?
| b | Water jar cleaned?
| c | Parts soaked in antiseptic
[5 empty] | Yes | No
| d | Use sterile (boiled tap) or distilled water? |  | Enter X in one box:
4 | Training - How many staff are trained for delivering oxygen therapies? |  |  | Enter # 0 if none | Respira-tory therapists | Nurses | Doctors | Other:
| a | Delivery oxygen by cannula or face mask
| b | How to titrate oxygen
| c | Humidifier set up and cleaning
| d | Infant nasopharyngeal O2 delivery
| e | Able to intubate
| f | Mechanical ventilator operation
| g | Operate high-flow nasal oxygen delivery set up
| h | IPC: Trained in airborne precautions (N95 +)
| i | IPC: Trained in droplet precautions

## P4 Inoperative Equipment
O2 Survey P4: Nonfunctioning equipment needing repair (light blue totals from P2)
WHY?
|  | TOTAL NOT FUNCTIONAL | No spare parts | No funds | No training for repair | No necessary parts | Not installed | No distributor | Single use and damaged | Other, specify
1 | Nasal cannula | 0
2 | Neonate/ Infant nasopharyngeal catheters | 0
3 | Face masks - single use | 0
4 | Face masks with reservoir bag | 0
5 | Bag Valve mask-adult | 0
6 | Bag Valve mask -child | 0
7 | Intubation sets | 0
8 | Oral airway | 0
9 | Nasopharengeal airway | 0
10 | Crico-thyroidotomy tray | 0
11 | Regulator/gauge- bullnose | 0
12 | Regulator/gauge- pin index | 0
13 | Flow splitter- # outlets | 0
14 | Mechanical ventilator - adult | 0
15 | Mechanical ventilator -child | 0
16 | Mechanical ventilator - portable | 0
17 | Bubble CPAP | 0
18 | CPAP | 0
19 | BiPAP | 0
21 | O2 terminal wall units - with valve, pressure & flow regulator | 0
22 | O2 terminal wall units without these valves and regulators | 0
23 | Oxygen tubing- sets | 0
24 | Patient monitor- integrated ECG | 0
25 | Patient monitor - no ECG | 0
26 | High flow nasal oxygen- child | 0
27 | High flow nasal oxygen- adult | 0
28 | Suction - manual | 0
29 | Suction - electric | 0
30 | Suction - central vacuum | 0
31 | Humidifier | 0
32 | Nebulizers | 0
33 | Oximeter - finger tip | 0
34 | Oximeter - hand held | 0
35 | Oximeter - table top | 0
36 | Probes for handheld, tabletop | 0
37 | Blood gas analyzer | 0
CONCENTRATORS
38 | 5 liter flow | 0
39 | 8 liter flow | 0
40 | 10 liter flow | 0
41 | > 10 liter flow | 0

## P5 Technician - UPS, generat
O2 Survey P5: Technician to complete details of power for facility
1 | 1a. WHAT IS THE SOURCE OF ELECTRICITY FOR THIS FACILITY? | Enter X in one box: | Central electricity grid | Power generator | Both | Other
1b. If Other, specify
| If Power generator or Both: | 1c. How many generators are available?
1d. What capacity are they in kVa?
2 | DOES EACH WARD/DEPARTMENT HAVE DEPENDABLE VOLTAGE STABILIZATION; DOUBLE CONVERSION UNINTERRUPTIBLE POWER SUPPLY? | Enter X in each box if present:
|  | Emergency room | Pediatric ward | Adult ward | Operating room | Maternity ward | ICU | Neonatal Unit | Lab | Radiology

## P6 Technician- central O2
O2 Survey P6: If piped central oxygen is available at this facility, a technician should complete this page
1 | If manifolds are used for multiple cylinders: | Manifold serial number | Location | Which wards/areas does this supply (if central, write 'all') | Number of cylinders manifold can fit | Type of switch (manual / automatic)
1
2
3
4
5
6
Further manifolds
2 | If bulk liquid oxygen tank is used: | 2a. What is the tank capacity in m3 (if other unit, indicate this)
2b. How often is the tank refilled? (e.g. weekly, monthly; if other, specify)
2c. How much does refilling cost, per m3?
3 | If there is an on-site PSA oxygen plant: | 3a. What is the manufacturer and model of the plant?
3b. What is the maximum production capacity of the plant? (in m3/hour, or if other units, indicate this)
3c. What is the average oxygen consumption per month (in m3, or if other units, indicate this)
3d. What is the percent concentration of oxygen produced by the plant? (%)
|  | Enter a X under Yes or No for the questions below |  |  | Yes | No
3e. Is the oxygen plant functional and operational?
3f. Is the oxygen plant operating 24 hours a day?
3g. Is there direct piping from the oxygen plant to wards/departments?
3h. Is the plant connected to backup electricity, like a generator?
3i. If yes, is this backup source dedicated to the oxygen plant alone?
3j. Is there a surge suppressor for the plant?
3k. Is there a voltage stabilizer for the plant?
3l. Is there a filling ramp or a manifold for filling cylinders?
3m. If Yes, how many bull nose cylinder connections does the ramp/manifold have?
3n. If Yes, how many pin index cylinder connections does the ramp/manifold have?
|  | 3p. Up to what pressure are the cylinders filled? (Enter answer under one unit only) |  | Bar | kPa | PSI
|  | 3q. Approximately how many cylinders of each type did the plant fill in the last month? |  |  | Number filled
Size C, 170 liters - pin index
Size D, 340 liters - pin index
Size E, 680 liters - pin index
Size F, 1360 liters - bullnose
Size G, 3400 liters - bullnose
Size H, 4100 liters - bullnose
Size J, 6800 liters - bullnose
Others (code, liters, connector type):

## P7 How many can be treated
O2 Survey P7: How many adults with respiratory distress can be managed on oxygen?
The data from the survey can be used to estimate how many adults in respiratory distress can be treated for one day, based on rate of O2, with the current oxygen sources in your hospital- if all cylinders full, all concentrators working or functional oxgyen plant. After cylinders used up, calculations for subsequent days rely on concentrators (or oxygen plant) plus refilled cylinders (factoring in time to refill).
1 |  | Liters/ min | Liters/day at this rate for one patient | Number of adults in wards plus emergency room | Total require-ment in liters/day
| a | 5 | 7200 |  | 0
| b | 8 | 11520 |  | 0
| c | 10 | 14400 |  | 0
| d | 15 | 21600 |  | 0
| e | 30 | 43200 |  | 0
2 |  |  |  |  | Patients | O2 liters per day | Liters per 10-day admission
| Total adults with COVID-19 in your district* |  |  |  |  | 0 | 0
| Estimated number with severe COVID-19 |  |  |  | 0 | 0 | 0
| Estimated number with critical COVID-19 |  |  |  | 0 | 0 | 0
*(NOTE total oxygen requirements in this row are a sum of the two rows below: severe + critical)
Capacity by each cylinder:

## P8 For ventilator-CPAP-BiPAP
P8 O2 Survey: Additional consumable equipment for mechanical ventilator, CPAP, BiPAP
Not all items are needed for each machine: consult your technician or the machine's manual
As these vary by machine, list type of machines below
1 | Enter number of consumable item (0 if none) If counting boxes or packs of items, enter numbers after multiplying by units per box: | Emergency room | Pediatric ward | Adult ward | Operating room | Maternity ward | ICU | Neonatal Unit | IN STORAGE | TOTALS
a | Filters for ventilators | [8 empty] | 0
b | Connectors for ventilator to ET tube | [8 empty] | 0
c | Nasal prongs for high flow nasal O2 | [8 empty] | 0
d | Face masks for CPAP/BiPAP | [8 empty] | 0
e | Nasal masks for CPAP/BiPAP | [8 empty] | 0
f | Nasal prongs for CPAP/BiPAP | [8 empty] | 0
g | Headgear for each design face/nasal mask (for CPAP/BIPAP) | [8 empty] | 0
h | CPAP/BiPAP hose | [8 empty] | 0
i | Connector from CPAP/BiPAP hose into machine with oxygen enrichment | [8 empty] | 0
j | Filters for CPAP/BiPAP | [8 empty] | 0
k | Connectors for ventilator to O2 supply | [8 empty] | 0
2 | Enter ADULT VENTILATOR brand & model, and for each, number present in each area | Emergency room | Pediatric ward | Adult ward | Operating room | Maternity ward | ICU | Neonatal Unit | IN STORAGE | TOTALS
0
0
3 | Enter CHILD VENTILATOR brand & model, and for each, number present in each area | Emergency room | Pediatric ward | Adult ward | Operating room | Maternity ward | ICU | Neonatal Unit | IN STORAGE | TOTALS
0
0
4 | Enter PORTABLE VENTILATOR brand & model, and for each, number present in each area | Emergency room | Pediatric ward | Adult ward | Operating room | Maternity ward | ICU | Neonatal Unit | IN STORAGE | TOTALS
0
0
5 | Enter CPAP machine brand & model, and for each, number present in each area | Emergency room | Pediatric ward | Adult ward | Operating room | Maternity ward | ICU | Neonatal Unit | IN STORAGE | TOTALS
0
0
6 | Enter BiPAP machine brand & model, and for each, number present in each area | Emergency room | Pediatric ward | Adult ward | Operating room | Maternity ward | ICU | Neonatal Unit | IN STORAGE | TOTALS
0
0